Heart Rate & Temperature

Monitor Simulation

Project Documentation - Group 15

KNUST  ·  Microprocessors ·  Computer Engineering

Table of Contents

1. Introduction	3

1.1 Project Objectives	3

1.2 Simulation Tool	3

2. Problem Statement	4

3. Methodology	5

3.1 Design Approach	5

3.2 Component Selection Rationale	5

4. System Components & Connections	6

4.1 Component Overview	6

4.2 Pin Connection Table	6

4.3 LCD Wiring Details	6

5. Arduino Code & Documentation	7

5.1 Libraries Used	7

5.2 Complete Code	7

5.3 Key Code Decisions	9

6. Oscilloscope Setup & Output	10

6.1 Connection	10

6.2 Recommended Settings	10

6.3 BPM Verification Formula	10

7. Alarm System	11

7.1 Thresholds	11

7.2 Buzzer Behaviour	11

8. Running the simulation	12

8.1 Heart Beat Sensor Hex File	12

8.2 Arduino Uno Hex File	12

9. Conclusion	13

9.1 Findings	13

9.2 Challenges Encountered	13

9.3 Possible Improvements	13

10. References	14

# 1. Introduction

This document presents the design, implementation, and simulation of a real-time Heart Rate and Temperature Monitoring System using Arduino Uno in Proteus 8 Professional. The system integrates a Heart Beat Sensor (HB1) and an LM35 temperature sensor to continuously measure and display patient vitals on a 16x2 LCD screen. A buzzer alarm activates automatically when the heart rate falls outside safe clinical thresholds, alerting to potential tachycardia or bradycardia conditions with a Heart Attack warning.

## 1.1 Project Objectives

- Simulate a dual-sensor health monitor in Proteus 8 Professional

- Display real-time BPM and temperature readings on a 16x2 LCD

- Trigger an audible buzzer alarm and visual warning when BPM exceeds 100 (tachycardia) or drops below 60 (bradycardia)

- Visualise the heartbeat waveform on a digital oscilloscope

- Produce clean, well-commented code that can be explained during presentation

## 1.2 Simulation Tool

Proteus 8 Professional was selected as the simulation environment due to its comprehensive AVR microcontroller support, built-in virtual instruments (oscilloscope, voltmeter), and compatibility with the Arduino Uno (ATmega328P) and the TheEngineeringProjects Heart Beat Sensor library.

# 2. Problem Statement

Manual monitoring of patient vitals is slow, error-prone, and resource-intensive. Patients with cardiac conditions require continuous automated tracking of both heart rate and body temperature to enable timely clinical intervention. Without an automated alert system, dangerous conditions such as tachycardia (heart rate above 100 BPM) and bradycardia (heart rate below 60 BPM) may go undetected, leading to life-threatening outcomes.

The core challenge addressed by this project is:

How can we design a low-cost embedded system that continuously monitors heart rate and body temperature, alerts caregivers to dangerous conditions, and displays real-time readings — all within a simulated environment?

# 3. Methodology

## 3.1 Design Approach

The project followed a five-stage design methodology:

- Component Selection — Identified compatible components available in the Proteus 8 library

- Schematic Design — Wired all components in Proteus Schematic Capture

- Code Development — Combined PulseSensor and LiquidCrystal libraries with alarm logic

- Simulation & Testing — Loaded compiled HEX, tuned oscilloscope, verified LCD output

- Alarm Integration — Implemented non-blocking buzzer using millis() timing

## 3.2 Component Selection Rationale

| Component | Selection Reason |
| --- | --- |
| Heart Beat Sensor (HB1) | Available in TEP Proteus library; runs internal ATmega328P firmware |
| LM35 Temperature Sensor | Linear 10mV/°C output; directly readable by Arduino ADC |
| Arduino Uno (ATmega328P) | Widely supported in Proteus; compatible with all required libraries |
| 16x2 LCD (LM016L) | Standard Proteus component; works with LiquidCrystal library |
| Active Buzzer | Simple digital output; ideal for alarm signalling on D8 |
| Digital Oscilloscope | Built-in Proteus virtual instrument; no extra library needed |

# 4. System Components & Connections

## 4.1 Component Overview

The system consists of six main components connected to an Arduino Uno microcontroller:

## 4.2 Pin Connection Table

| Signal / Component | Arduino Pin | Notes |
| --- | --- | --- |
| Heart Beat Sensor (HB1) | Analog A0 | VccOUT pin of HB1 → A0 |
| LM35 Temperature Sensor | Analog A1 | VOUT via R1 resistor → A1 |
| Onboard LED | Digital D13 | Blinks with each detected heartbeat |
| Buzzer (alarm) | Digital D8 | Active HIGH; sounds during alarm |
| LCD RS (Register Select) | Digital D7 | 4-bit LCD control |
| LCD E (Enable) | Digital D6 | 4-bit LCD control |
| LCD D4 | Digital D5 | 4-bit data bus |
| LCD D5 | Digital D4 | 4-bit data bus |
| LCD D6 | Digital D3 | 4-bit data bus |
| LCD D7 | Digital D2 | 4-bit data bus |
| Oscilloscope Ch A | Junction A0/HB1 | Tapped directly at signal junction |

## 4.3 LCD Wiring Details

The 16x2 LCD (LM016L in Proteus) is wired in 4-bit mode to conserve Arduino pins:

| LCD Pin | Name | Connect To |
| --- | --- | --- |
| 1 | VSS | GND |
| 2 | VDD | +5V |
| 3 | VEE (Contrast) | GND — gives full contrast in simulation |
| 4 | RS | Arduino D7 |
| 5 | RW | GND — write-only mode |
| 6 | E (Enable) | Arduino D6 |
| 11–14 | D4–D7 | Arduino D5–D2 |
| 15 | BLA (Backlight +) | +5V |
| 16 | BLK (Backlight −) | GND |

# 5. Arduino Code & Documentation

## 5.1 Libraries Used

| Library | Purpose |
| --- | --- |
| PulseSensorPlayground | Handles heart rate detection, BPM calculation, and beat detection logic |
| LiquidCrystal | Drives the 16x2 LCD in 4-bit mode |

## 5.2 Complete Code

/*

* Combined Heart Rate + LM35 Temperature Monitor

* - PulseSensor on Analog Pin A0

* - LM35 on Analog Pin A1

* - 16x2 LCD in 4-bit mode

* - LED blink on Pin 13 with heartbeat

* - Buzzer on Pin 8 for Tachycardia / Bradycardia alarm

*/

#define USE_ARDUINO_INTERRUPTS false

#include <PulseSensorPlayground.h>

#include <LiquidCrystal.h>

// Pin Definitions

const int PULSE_PIN   = 0;   // PulseSensor  -> A0

const int TEMP_PIN    = 1;   // LM35         -> A1

const int LED_PIN     = 13;  // Onboard LED

const int BUZZER_PIN  = 8;   // Buzzer       -> D8

// Alarm Thresholds

const int TACHY_THRESHOLD = 100;  // Above 100 BPM = Tachycardia

const int BRADY_THRESHOLD = 60;   // Below  60 BPM = Bradycardia

// Buzzer Timing (non-blocking)

const unsigned long BEEP_INTERVAL = 400;

const unsigned long BEEP_DURATION = 150;

unsigned long lastBuzzerBeep = 0;

bool buzzerState = false;

int Threshold = 550;

PulseSensorPlayground pulseSensor;

LiquidCrystal lcd(7, 6, 5, 4, 3, 2);

int currentBPM  = 0;

unsigned long lastLCDUpdate = 0;

const unsigned long LCD_INTERVAL = 1000;

byte heartChar[8] = {

0b00000, 0b01010, 0b11111, 0b11111,

0b01110, 0b00100, 0b00000, 0b00000

};

byte bellChar[8] = {

0b00100, 0b01110, 0b01110, 0b01110,

0b11111, 0b00000, 0b00100, 0b00000

};

void setup() {

Serial.begin(9600);

pinMode(BUZZER_PIN, OUTPUT);

digitalWrite(BUZZER_PIN, LOW);

lcd.begin(16, 2);

lcd.createChar(0, heartChar);

lcd.createChar(1, bellChar);

lcd.setCursor(0, 0); lcd.print('  Health Monitor');

lcd.setCursor(0, 1); lcd.print('  Initializing..');

delay(2000); lcd.clear();

lcd.setCursor(0, 0); lcd.print('BPM:');

lcd.setCursor(0, 1); lcd.print('Temp:');

pulseSensor.analogInput(PULSE_PIN);

pulseSensor.blinkOnPulse(LED_PIN);

pulseSensor.setThreshold(Threshold);

if (pulseSensor.begin()) Serial.println('PulseSensor started OK.');

}

void loop() {

// STEP 1: Read Heart Rate

if (pulseSensor.sawStartOfBeat()) {

currentBPM = pulseSensor.getBeatsPerMinute();

Serial.print('BPM: '); Serial.println(currentBPM);

}

// STEP 2: Read Temperature (LM35: 10mV/degC)

int rawTemp   = analogRead(TEMP_PIN);

float voltage = rawTemp * (5000.0 / 1024.0);

float tempC   = voltage / 10.0;

// STEP 3 & 4: Non-blocking buzzer alarm

unsigned long now = millis();

bool alarmActive = (currentBPM > TACHY_THRESHOLD ||

(currentBPM > 0 && currentBPM < BRADY_THRESHOLD));

if (alarmActive) {

if (!buzzerState && (now - lastBuzzerBeep >= BEEP_INTERVAL)) {

digitalWrite(BUZZER_PIN, HIGH);

buzzerState = true; lastBuzzerBeep = now;

}

if (buzzerState && (now - lastBuzzerBeep >= BEEP_DURATION)) {

digitalWrite(BUZZER_PIN, LOW); buzzerState = false;

}

} else {

digitalWrite(BUZZER_PIN, LOW); buzzerState = false;

}

// STEP 5: Update LCD every 1 second

if (now - lastLCDUpdate >= LCD_INTERVAL) {

lastLCDUpdate = now;

if (alarmActive) {

lcd.setCursor(0, 0); lcd.write(byte(1));

lcd.print('HEART ATTACK!'); lcd.write(byte(1));

lcd.setCursor(0, 1); lcd.print('BPM:');

lcd.print(currentBPM);

lcd.print(currentBPM > TACHY_THRESHOLD ? ' HIGH  ' : ' LOW   ');

} else {

lcd.setCursor(4, 0);

if (currentBPM > 0) {

lcd.print(currentBPM); lcd.print(' ');

lcd.write(byte(0)); lcd.print('  ');

} else { lcd.print('---   '); }

lcd.setCursor(5, 1);

lcd.print(tempC, 1); lcd.print((char)223); lcd.print('C  ');

}

Serial.print('Temp: '); Serial.print(tempC, 1); Serial.println(' C');

}

delay(20);

}

## 5.3 Key Code Decisions

| Decision | Explanation |
| --- | --- |
| USE_ARDUINO_INTERRUPTS false | Proteus does not support low-level AVR timer interrupts reliably; false mode uses polling instead |
| Non-blocking buzzer with millis() | Using delay() for the buzzer would pause heart rate reading during an alarm — millis() prevents this |
| LCD_INTERVAL = 1000ms | Refreshing faster causes LCD flicker in Proteus simulation; 1 second gives stable readable output |
| tempC = voltage / 10.0 | LM35 outputs exactly 10mV per degree Celsius; dividing millivolts by 10 gives Celsius directly |
| Threshold = 550 | Default PulseSensor beat detection threshold; prevents noise from triggering false beats |

# 6. Oscilloscope Setup & Output

## 6.1 Connection

The oscilloscope Channel A is connected at the junction between the Heart Beat Sensor VccOUT pin and Arduino Analog Pin A0. The oscilloscope GND terminal connects to the schematic GND. Channels B, C, and D are left disconnected.

## 6.2 Recommended Settings

| Control | Setting | Reason |
| --- | --- | --- |
| Time/Div | 200 mS | Heartbeat at 75 BPM has an 800ms period; needs wide time window |
| Channel A Volts/Div | 1 V | Signal swings 0–3V; 1V/div keeps waveform within visible range |
| Channel A Coupling | AC | Removes DC offset; centres waveform on display midline |
| Trigger Source | Channel A | Triggers oscilloscope sweep on the heartbeat signal itself |
| Trigger Mode | Auto | Keeps display stable even between beats |
| Channels B, C, D | OFF | Removes flat baseline lines from unused channels |

## 6.3 BPM Verification Formula

The oscilloscope can be used to manually verify the BPM reading shown on the LCD:

BPM = 60,000 ÷ period (ms)

For example, if peaks are 800ms apart: BPM = 60,000 ÷ 800 = 75 BPM. This cross-checks the PulseSensorPlayground library output displayed on the LCD.

# 7. Alarm System

## 7.1 Thresholds

| Condition | BPM Range | LCD Row 0 | LCD Row 1 | Buzzer |
| --- | --- | --- | --- | --- |
| Normal | 60 – 100 BPM | BPM: XX heart | Temp: XX.X°C | Silent |
| Tachycardia | > 100 BPM | bell HEART ATTACK! bell | BPM:XX HIGH | Beeping |
| Bradycardia | < 60 BPM | bell HEART ATTACK! bell | BPM:XX LOW | Beeping |

## 7.2 Buzzer Behaviour

The buzzer operates in a non-blocking pattern using millis() timing:

- Beep ON for 150ms

- Beep OFF for 250ms (400ms interval minus 150ms duration)

- Pattern repeats continuously while alarm condition persists

- Immediately silences when BPM returns to the normal 60–100 range

# 8. Running the simulation

The proteus simulation file is included in the folder and named as “heart_rate_monitor_and_temperature_sensor buzz.pdsprj”.

## 8.1 Heart Beat Sensor Hex File

The heart beat sensor from theengineeringprojects.com needs a “. hex” file, provided by the authors, in order to operate properly. The “. hex” file is included in this folder named as “HeartBeatSensorTEP2.HEX”. To add the file to the sensor component in your proteus simulation:

- Double-click the heart beat sensor.

- Click the folder icon next to the “Program File” field.

- Navigate to the “. hex” location and select it.

- Click “ok”.

## 8.2 Arduino Uno Hex File

In order to virtually upload the Arduino code to the simulated microcontroller, a “. hex” must be generated from the written program and added to the simulation in a similar fashion to the steps above. The “. hex” file for this is also included in this folder named as “heart_rate_temp.ino.hex”. To add the file to the microcontroller component in your proteus simulation:

- Double-click the microcontroller.

- Click the folder icon next to the “Upload Hex File” field.

- Navigate to the “. hex” location and select it.

- Click “ok”.

That is all the set up done now. You can click the run button at the bottom left of the window.

# 9. Conclusion

## 9.1 Findings

The simulation successfully demonstrated a functional real-time health monitoring system. The Arduino Uno processed readings from both the Heart Beat Sensor and LM35 simultaneously, updating the LCD every second with BPM and temperature values. The buzzer alarm correctly triggered at the set thresholds, and the oscilloscope displayed a repeating heartbeat waveform consistent with the BPM output.

## 9.2 Challenges Encountered

- The Heart Beat Sensor (HB1) from TheEngineeringProjects runs an internal ATmega328P firmware — it is not a simple analog sensor. This means BPM cannot be changed dynamically during simulation.

- USE_ARDUINO_INTERRUPTS must be set to false in Proteus because the simulator does not support low-level AVR hardware interrupts used by the PulseSensor library.

- The oscilloscope horizontal time base required careful tuning — the value box interprets typed input in seconds, so entering 200ms requires typing 0.2 (not 200).

- Filter chain attempts (LPF → Op-Amp → HPF) intended to shape the waveform instead caused phase shift, doubling the apparent BPM reading to ~230 BPM.

## 9.3 Possible Improvements

- Replace the HB1 with a real analog signal generator (sine wave, 1.25 Hz, 1.5V amplitude, 1.5V offset) for a more realistic PPG waveform on the oscilloscope.

- Add a Wi-Fi or Bluetooth module (ESP8266/HC-05) for remote patient monitoring and data logging.

- Include an SpO2 sensor (MAX30102) for blood oxygen saturation measurement alongside BPM.

- Replace the buzzer with a multi-tone alarm system that uses different tones for tachycardia vs bradycardia.

- Add a real-time clock (RTC) module to timestamp readings for patient record keeping.

# 10. References

- TheEngineeringProjects.com — Heart Beat Sensor Proteus Library (HeartBeatSensorTEP.LIB)

- Arduino.cc — PulseSensorPlayground Library Documentation

- Arduino.cc — LiquidCrystal Library Documentation

- Proteus 8 Professional — Labcenter Electronics

- Texas Instruments — LM35 Precision Centigrade Temperature Sensor Datasheet

- Atmel — ATmega328P Datasheet (Arduino Uno Microcontroller)

# Group Members

- Emzera Akosua Ayesu  7267123

- Claudette Cobbah - 7268823

- Amissah-Ayiku Daniel Dromo - 7262923

- Asiseh Sesinam Akosua 7265923

- Asaah Ayebea Ayeh - 7267023

- Asare-Boateng Kwasi Boateng- 7265623

- Allan Fiifi Buaful - 7268723

- Nana Kwame Amoah Danquah – 7269523

- Donkor Joshua Akwasi - 7270123

- Kwabena Boateng Gyau Baffour 7272123